Please leave these first three lines blank for the Editors

Please leave these first three lines blank for the Editors

Please leave these first three lines blank for the Editors

A MECHANICAL ANALYSIS OF POLYPROPYLENE PROSTHESES USED IN PELVIC FLOOR REPAIR

Please leave this line blank

Pedro Martins1(*), José Afonso2, Renato Jorge3, António Ferreira4, M. Girao5, Teresa Mascarenhas6, António Fernandes7, J. Bernardes8, E. Bacarat9, G.Rodrigues de Lima10, and Belmiro Patricio11

1, 3, 7 Institute of Mechanical Engineering (IDMEC), University of Porto, Porto, Portugal

2, 5, 9, 10 Department of Gynecology, Federal University of São Paulo, São Paulo, Brazil

4 Department of Mechanical Engineering and Industrial Management (DEMEGI), University of Porto, Portugal

3, 8, 11 Department of Gynecology, Faculty of Medicine, University of Porto, Porto, Portugal

(*) Email: palsm@fe.up.pt

Please leave these two lines blank

Please leave these two lines blank

ABSTRACT

This work compares the tensile (uniaxial tension test) stiffness of different meshes used as prostheses in pelvic floor repair surgeries. Five different commercial meshes were analysed, all made from polypropylene but with different geometries: Aris™, TVTO™, Uretex™, Avaulta™ and Auto Suture ™. The mechanical tests performed allowed to identify significant differences on the mechanical properties of the polypropylene meshes used in urogynecology surgeries.

Please leave this line blank

INTRODUCTION

There is erosion phenomena associated with polypropylene prostheses used in pelvic floor repair surgeries. This phenomenon consists on the inflammation of the biologic structures directly in contact with the artificial material. Several problems can occur as a consequence from pain and discomfort to an extreme loss of tissue, leading ultimately to a new surgery for prosthesis removal or substitution. For the case of stress urinary incontinence (SUI), the erosion rate associated is approximately 1–3% (Petros, 2007). Some evidence points to the fact that mesh stiffness may be related with erosion (Siegel, 2005). The flexibility differences of the meshes (all monofilament polypropylene tapes with macropores) are due to the differences in mesh weave and surgical suture thickness.

A total of 15 individual tests were performed. The sample length was 40 mm (±0.5 mm) with slightly varying width according with their technical specifications (≈10 mm). The mechanical tests were performed at slow displacement rate (5 mm/s) and at room temperature (≈25º C). The mechanical testing was carried out on a testing setup purposely built for low load soft tissue testing (Martins, 2006). According with a previous study (Dietz, 2003), temperature doesn’t influence the results significantly. From the tensile test two stiffness parameters were extracted EI and EII. EI used load-displacement data in the range of 0-500 g, the estimated physiologic load range (Dietz, 2003). For the calculation of EII the load range was 500-2000 g.

Please leave this line blank

RESULTS AND CONCLUSIONS

The results from the tensile tests are shown in Fig. 1. The load-displacement curve has two different regions. The first region is nonlinear and evolves to an approximately linear region. Table 1 shows the results for the stiffness parameters EI and EII.

The most significant difference of tensile stiffness behaviour appears between Aris™ and TVTO™ on the EI (p=0.001) and EII (p=0.0003) regions.

Table 1 Uniaxial tension test results

| Tapes | EI [(N/m)x1000] | StD | EII [(N/m)x1000] | StD |
| --- | --- | --- | --- | --- |
| Aris™ | 2.3898 | 0.2154 | 5.3705 | 0.5432 |
| TVTO™ | 0.9428 | 0.1919 | 1.3083 | 0.2075 |
| Uretex™ | 1.1367 | 0.1918 | 3.2231 | 0.1853 |
| Avaulta™ | 1.7986 | 0.1925 | 3.7810 | 0.7626 |
| Auto Suture™ | 1.0507 | 0.0766 | 2.2624 | 0.2658 |

Fig.1 Tensile test results

This study shows that there are substantial differences on the mechanical properties of different urogynecology meshes. Further tests should be performed in order to analyze other mechanical properties, such as flexural properties.

Please leave this line blank

ACKNOWLEDGMENTS

The authors gratefully acknowledge the funding by Ministério da Ciência, Tecnologia e Ensino Superior, FCT, Portugal, under grants PTDC/SAU-BEB/71459/2006 and SFRH/BD/41841/2007.

Please leave this line blank

REFERENCES

[1] Dietz H, Vancaillie P, Svehla M, Walsh W, Steensma A, Vancaillie T. Mechanical properties of urogynecologic implant materials. International Urogynecology J, 2003, 14, p. 239-243.

[2] Martins PALS, Jorge RMN, Ferreira AJM. A comparative study of several material models for prediction of hyperelastic properties: Application to Silicone-Rubber and Soft Tissues. Strain, 2006, 42, p. 135-147.

Please use numbered brackets [#] to separate your references.